# Maintaining ExoPrep Methods

By Kevin Haynie

10/10/2025

# Importing ExoPrep Methods

The ExoPrep GUI uses methods to determine which steps to send to the ExoPrep instrument.  The contents of these methods is driven from \*.method files which have a specific format as defined below.  New methods can be imported and existing ones can be updated.  Existing methods can be deleted by simply deleting the file.

Method files are stored in the user's `\ProgramData\ExoPrep\methods` directory and the name of the method is defined by the filename prefix.  For example, `c:\programdata\ExoPrep\PCR.method` will be called `PCR` within the ExoPrep GUI.

All default methods are stored in the `\ProgramData\ExoPrep\methods\default` directory.

All method files use YAML syntax and must be syntactically correct YAML.  Users should validate that their files are syntactically correct YAML before trying to import into ExoPrep GUI.  Be careful to use proper indentation since YAML depends on these indentations to know hierarchy of objects.

There are 3 main sections in all *.method files: `Method` and `Steps` and `Config`. Each has its own specific format as defined below.  Comments can be included in these files by prefixing lines with the hashtag character (#). Blank lines are acceptable as well.

## Method section

The following properties can be set in the `[Method]` section:

| Property Name | Description | Example | Default |
| --- | --- | --- | --- |
| author | Name of the person or company who created this method | author: John Brown | -none- |
| version | Version number of this method | version: 1.0 | 1.0 |
| description | user defined description of what this method is intended for (e.g. downstream applications) | description : Method used for Mass Spectometry | -none- |
| type | Type of method (standard, calibration, selfTest); only standard methods will show in the methods list on the Select Method page. Only calibration type methods will show in Tools¦Calibration. Only selfTest type methods will show in the Tools¦Self Test page. | type: standard | standard |
| shared | Set to true if this method, when loaded, is to be loaded as a shared method where all users in the system will be able to view and run. Note: All default methods installed by the ExoPrepGUI installer are set to be shared. | Shared: true | false |
| filmRequired | the biochip film required by this method (massSpec, immunochemistry, PCR, any). Use `all` to allow any film to be used with this method | filmRequired: massSpec | all |
| instruments | List of Exokeryx instruments required for this method (default is ‘all’). Use `all` to allow any instrument to be used with this method. Available instruments are: - ExoDemo - ExoPrepFocaccia - ExoPrep4UP - DemeterEVPrep - ExoPrepEmulator | instruments: [DemeterEVPrep, ExoPrepEmulator] | all |
| cartridges | List of cartridge part numbers which this method requires. ‘*’ may be used within a part number for easier matching. Set to `all` to allow any type of cartridge to be used with this method | cartridges: [12345,1112*] | all |
| minSampleVolume | minimum sample required in uL | minSampleVolume: 100 | 100 |
| minEluateVolume | Minimum eluate volume required in uL | minEluateVolume: 50 | 50 |
| minSyringeVolume | Minimum syringe volume/size required for this method to run properly | minSyringeVolume: 300 | 300 |
| manualMode | `true` to disable instrument reset and priming | manualMode: true | false |
| constants | list of values which can be used within method parameters. Each constant has a name and a value. | constants: - DEPVoltage: 11 - DEPFreq: 2703 | -none- |

## Steps section

The `Steps` defines a list of actions which will be taken on the instrument for a particular method.  Each step has an action, and each action has its own set of parameters as defined by the following table.

All steps require a name parameter which defines the name of the step. duration can be included as a parameter to run the action for a preset amount of seconds.

| Action | Purpose and Parameters | Description/Parameter Values | Default | Required |
| --- | --- | --- | --- | --- |
| circulateFluidInChip (rev 2) |  | Moves pipette to docking well, pulls .05uL from the docking well to form negative pressure, opens the pinch valve, then swishes fluid currently stored in the processing well into and out of the chip and finally into the docking well. Duration parameter (default of 120 secs) must be long enough to complete all swish cycles or this action will timeout. - circulateFluidInChip: { name: swish 5uL of sample over the chip at 1uL/min using default pause time, volume: 5.0, speed: 1.0, numCycles: 2, duration: 604 } - circulateFluidInChip: { name: swish sample over the chip for total of 6 minutes at 2 uL/min, pulling/pushing for 30 seconds and delaying 5 seconds between directions, duration: 420, volume: 1.0, speed: 2.0, numCycles: 12, pauseTime: 5 } |  |  |
|  | volume | volume in uL to pull into the flow cell and over the chip (decimal OK) |  | Yes |
|  | speed | speed in uL/min to pull/push while swishing (decimal OK); if pullSpeed is included, this parameter represents the push speed. |  | Yes |
|  | numCycles | number of push/pull movements to run |  | Yes |
|  | pauseTime | number of seconds to wait between each direction change | 2 | No |
|  | pullSpeed | Speed in uL/min when pulling fluid |  | No |
|  | sourceWell | outlet, tips, reagents, waste, inlet, processing, docking, endzone |  | No |
|  | sourceNumber | well number (1-4) | 1 | No |
|  | sourceHeightOffset | Low – apply a low Z height offset Medium – apply a medium Z height offset High – apply a high Z height offset None – do not apply any offset | None | No |
| dropTip (rev 2) |  | Return current tip on pipette back to the cartridge. No parameters needed for this action. - dropTip: { name: drop current tip back into cartridge well } |  |  |
| emptyTip (rev 2) |  | Empty the content of the tip to specified well. - emptyTip: { name: clean out any remaining fluid in pipette tip and put into waste well by default } - emptyTip: { name: empty fluid from pipette into output well, destWell: outlet } - emptyTip: { name: empty fluid from pipette into reagent port 2, destWell: reagent, destNumber: 2 } |  |  |
|  | destWell | outlet, tips, reagents, waste, inlet, processing, docking, endzone | waste | No |
|  | destNumber | well number (1-4) | 1 | No |
|  | tapCount | Number of times to tap on the destination well after emptying; if not specified, the tap count will be set to the tap count specified in exoprep.json. If tap count does not exist in json file, the value hard coded in exoPrep API will be used. | Default exoprep.json value or default API value | No |
| getDepConfig | -none- | Prompt user for a list of sockets which will be used in future setDEP actions which have socket set to config. This can be used to subset which sockets will be energized. - getDepConfig: { name: prompt user to select set of sockets which will be energized in future setDEP actions } |  |  |
| getDoorReport | -none- | Obtain a snapshot of the status of the instrument door - getDoorReport: { name: get a snapshot of the instrument door status } |  |  |
| getEvent |  | Obtain a snapshot report of a statistic for a particular statistic - getEvent: { name: get a snapshot of temperatures for all sockets, event: temperature } |  |  |
|  | event | current, temperature, pressure, voltage, frequency, tip, all | all | No |
| getPressureReport | -none- | Obtain a snapshot of the status of all chip pressures - getPressureReport: { name: Refresh status of chip pressures } |  |  |
| getTECReport | -none- | Obtain a snapshot of the status of all TEC modules - getTECReport: { name: Refresh status of TEC modules } |  |  |
| logState |  | Append a timestamped ISO8601 message to the current exoPrep\logs\state-<run-name>-<method-name>.csv file. If message is not provided, the name of the step will be used as the message. - logState: { name: “SLOT, DEPVPP, DEPFREQ, TEMPSET, TEMPAUTO”, includeTimestamp: false } - logState: { name: “record header”, includeTimestamp: false, message: “SLOT, DEPVPP, DEPFREQ, TEMPSET, TEMPAUTO” } - logState: { name: “1, 10, 30000, 15, 1” } |  |  |
|  | message | String to append to the state-* file |  | No |
|  | includeTimestamp | True to always prepend the current date/time before the message; false to exclude | True | No |
| moveFluid (rev 2) |  | Move fluid between cartridge wells. This includes pulling fluid from a source cartridge well, moving pipette to a destination well location, and then pushing that fluid into that destination well. The speed of pull and push can be configured via the pullSpeed and pushSpeed parameters. The amount of fluid to move is driven by the volume parameter. A tip must be present before this action can take place. The source and destination wells must be different. - moveFluid: { name: move 10uL of sample to processing well, sourceWell: inlet, destWell: processing, volume: 10, pullSpeed: 2.0, pushSpeed: 2.0 } - moveFluid: { name: move 10uL of reagent 2 to processing well using a low reagent 2 offset, sourceWell: reagents, sourceNumber: 2, sourceOffset: low, destWell: processing, volume: 10, pullSpeed: 5.0, pushSpeed: 3.0 } - moveFluid: { name: pull 4 uL of primer through chip and into waste, sourceWell: docking, destWell: waste, volume: 4, pullSpeed: 3.0, pushSpeed: 5.0 } |  |  |
|  | sourceWell | home, outlet, tips, reagents, waste, inlet, processing, docking, endzone |  | Yes |
|  | sourceNumber | well number (1-4) | 1 | No |
|  | sourceHeightOffset | Low – apply a low Z height offset Medium – apply a medium Z height offset High – apply a high Z height offset None – do not apply any offset | None | No |
|  | destWell | home, outlet, tips, reagents, waste, inlet, processing, docking, endzone |  | Yes |
|  | destNumber | well number (1-4) | 1 | No |
|  | destHeightOffset | Low – apply a low Z height offset Medium – apply a medium Z height offset High – apply a high Z height offset None – do not apply any offset | None | No |
|  | volume | volume in uL of fluid to transfer (decimal values allowed) |  | Yes |
|  | pullSpeed | flow rate in uL/min when pulling from source well (decimal values allowed) |  | Yes |
|  | pushSpeed | flow rate in uL/min when pushing to destination well (decimal values allowed) |  | Yes |
|  | batchVolume | Batch volume in uL. Set to none to prevent batching. Set to auto to have the instrument calculate the optimal batch volume. At least 100uL of volume must be specified for batching to take place. | None | No |
| pickupTip (rev 2) |  | Pickup a pipette tip and drop current one if necessary - pickupTip: { name: pickup tip 1, number: 1 } - pickupTip: { name: pickup tip 2, number: 2 } |  |  |
|  | number | optional index within position (base 1) | 1 | No |
| playSound |  | Play an audio file to the default speaker - playSound: { name: play tada 2 times, filename: tada.mp3, numTimes: 2 } |  |  |
|  | filename | The filename of the audio file to play. If no path is included, file will be played from the \ExoPrep\audio directory. |  | Yes |
|  | numTimes | total number of times to play this audio file | 1 | No |
| pullFluidThroughChip (rev 2) |  | Pulls fluid from the specified sourceWell through the flow cell. If sourceWell is not specified, will always pull from the docking well. The pinch valve will be opened during the pull and then closed after the pull. - pullFluidThroughChip: { name: pull 5 uL of sample over the chip at 1uL/min, volume: 5.0, speed: 1.0 } |  |  |
|  | volume | volume in uL to pull into the flow cell and over the chip (decimal OK) |  | Yes |
|  | speed | pull flow rate in uL/min (decimal OK) |  | Yes |
|  | sourceWell | outlet, tips, reagents, waste, inlet, processing, docking, endzone |  | No |
|  | sourceNumber | well number (1-4) | 1 | No |
|  | sourceHeightOffset | Low – apply a low Z height offset Medium – apply a medium Z height offset High – apply a high Z height offset None – do not apply any offset | None | No |
|  | runInBackground | Start pull and return immediately; this can be used to start a pull and return immediately so that setDEP or other actions can be executed right away. To know when this pull is finished, use waitForBackgroundAction action. | False | No |
| repeatSteps |  | Repeat subsequent steps - repeatSteps: { name: Repeat next 2 steps 3 times, numSteps: 2, numLoops: 3 } |  |  |
|  | numSteps | The number of steps following this repeatSteps action which should be repeated |  | Yes |
|  | numLoops | The number of times those steps should be repeated |  | Yes |
| reset |  | Reset the instrument to a known state - reset: { name: recalibrate where home is, resetType: calibrateHome } |  |  |
|  | resetType | stop – DEPs off, release pinch valve calibrateHome – redefine home position, move to new home position goHome – move to last known home position recover – calibrate motors, empty fluid to waste, drop tips unclamp – move the gantry out of the way and unclamp |  | Yes |
| rinseWell (rev 2) |  | Move fluid from sourceWell to destWell using the pipette tip in use at the time this action is called. Then move fluid from destWell to waste well with either the existing tip or an optionally specified tip at tipNumber. - rinseWell: { name: clean our processing well before moving new fluid into it, sourceWell: reagents, sourceNumber: 3, destWell: processing, volume: 50, pullSpeed: 300, pushSpeed: 300, tipNumber: 2 } |  |  |
|  | tipNumber | optional tip to use when moving fluid from sourceWell and destWell (base 1) | Current Tip | No |
|  | sourceWell | home, outlet, tips, reagents, waste, inlet, processing, docking, endzone |  | Yes |
|  | sourceNumber | well number (1-4) | 1 | No |
|  | sourceHeightOffset | Low – apply a low Z height offset Medium – apply a medium Z height offset High – apply a high Z height offset None – do not apply any offset | None | No |
|  | destWell | home, outlet, tips, reagents, waste, inlet, processing, docking, endzone |  | Yes |
|  | destNumber | well number (1-4) | 1 | No |
|  | destHeightOffset | Low – apply a low Z height offset Medium – apply a medium Z height offset High – apply a high Z height offset None – do not apply any offset | None | No |
|  | volume | volume in uL of fluid to transfer (decimal values allowed) |  | Yes |
|  | pullSpeed | flow rate in uL/min when pulling from source well (decimal values allowed) |  | Yes |
|  | pushSpeed | flow rate in uL/min when pushing to destination well (decimal values allowed) |  | Yes |
| sendEmail |  | Sends an email to the email address of the currently logged in user using the SMTP settings defined in the GUI. The From address will be `no-reply@exokeryx.com` and the from name will be the name of the connected instrument. - sendEmail: { name: notify user instrument is ready, subject: instrument ready, body: Exokeryx instrument is now ready } |  |  |
|  | subject | Subject line of email |  | Yes |
|  | body | Plain text to send as email body |  | Yes |
| setClamp (rev 2) |  | Set state of cartridge clamp - setClamp: { name: clamp all cartridges, state: clamp } - setClamp: { name: release all cartridges, state: release } |  |  |
|  | state | clamp, release |  | Yes |
| setDEP |  | Turns on or off the DEP circuit given a frequency and voltage. DEP will not auto shut off until voltage set to 0. - setDEP: { name: turn all sockets DEP on at 11000 volts at 2700Hz, socket: all, voltage: 11000, frequency: 2700 } |  |  |
|  | socket | DEP socket (0-7), all, config Note: config requires that getDEPConfig has been run prior to this setDEP action | all | No |
|  | voltage | Voltage in mV to drive DEP(s) at (0 => turn DEP(s) off, min=5000, max=14000) |  | Yes |
|  | frequency | Frequency in Hz to drive DEP(s) at |  | Yes |
|  | pwmTimeOn | Number of seconds to cycle DEP in ON state (pwm cycling is enabled only if pwmTimeOff specified also) | ON full time | No |
|  | pwmTimeOff | Number of seconds to cycle DEP in OFF state (pwm cycling is enabled only if pwmTimeOn specified also) | Never OFF | No |
|  | pwmOffVoltage | Voltage in mV to drive DEP(s) at when cycled in OFF state (only used if pwm is enabled) | 0 | No |
|  | pwmOffFrequency | Frequency in Hz to drive DEP(s) at when cycled in the OFF state (only used if pwm is enabled on alpha or later instruments) | 0 | No |
| setLED |  | Sets the colors of various LEDs. Set color values to all zero to turn LED off. - setLED: { name: set all socket LEDs to blue, position: socketAll, color: “#00007f” } - setLED: { name: turn socket 1 (2nd from left) red, position: socket1, color: “#7f0000” } |  |  |
|  | position | socketAll; LEDs above all cartridge sockets on the front of the instrument; Yes. socket0; LED above leftmost cartridge socket on the instrument. socket1; LED above 2nd cartridge socket on the left of the instrument. socket2, socket3, …, socket7; LED above each one of the cartridge sockets. boardPrime; LED mounted on ‘prime’ board inside of instrument. boardTendrilB; LED mounted on ‘tendril B’ board inside of instrument. boardTendrilC; LED mounted on ‘tendril C’ board inside of instrument. reserved1; Reserved for future use. clearAll; Will turn OFF all LEDs on instrument; color parameter ignored. |
|  | color | RGB (red, green, blue) value in hex format preceded by a pound sign. FF indicates full intensity and 00 indicates no intensity. |  | Yes |
| setEventTimer |  | Starts periodic event reporting of one or more instrument statistics. Available statistics are: temperature, current, TEC, pressure, voltage, frequency - setEventTimer: { name: start event timer for all events every 3 seconds, stats: all, frequency: 3 } - setEventTimer: { name: start event timer for current events every 5 seconds, stats: current, frequency: 5} |  |  |
|  | stats | Which statistic(s) to report temperature – temperature of DEP chips current – current in mV being seen on DEP chip TEC – thermoelectric cooler status pressure – chip pressure status voltage – chip voltage status frequency – dep frequency status all – both temperature and current | all | No |
|  | frequency | How often to send statistics reports in seconds (0 = stop sending reports) | all | No |
| setMotorPosition (rev 2) |  | Diagnostics Only: sets various elements of instrument to specific positions |  |  |
|  | blocking | true – all motors will move at same time false – motors will move independently |  | Yes |
|  | relative | true – distances supplied in this action are relative to last known position false – distances supplied in this action are absolute and not dependent on previous position |  | Yes |
|  | fullStop | true – all motors will be stopped and other parameters on this action will be ignored false – motors will continue to run according to the parameters on this action |  | Yes |
|  | motorX | x gantry position in mm |  | Yes |
|  | motorZPipette | pipette lift position in mm |  | Yes |
|  | motorZClamp | cartridge clamp movement in mm |  | Yes |
|  | motorPinch | pinch valve movement in mm |  | Yes |
|  | motorStripTip | movement to detach the pipette tip in mm |  | Yes |
|  | motorPump | movement of pipette pump in mm |  | Yes |
| setMotorSpeed (rev 2) |  | Diagnostics Only: sets speed of various instrument motors |  |  |
|  | motorX | x gantry speed in mm/sec |  | Yes |
|  | motorZPipette | pipette lift speed in mm/sec |  | Yes |
|  | motorZClamp | speed in mm/sec of z movements of cartridge clamp |  | Yes |
|  | motorPinch | speed of pinch valve movement in mm/sec |  | Yes |
|  | motorStripTip | speed of pipette detachment in mm/sec |  | Yes |
|  | motorPump | speed of movement of pipette pump in mm/sec |  | Yes |
| setPinchValve (rev 2) |  | Opens/closes pinch valves under processing well - setPinchValve: { name: close pinch valves, state: close } - setPinchValve: { name: open pinch valves after processing well is full, state: open } |  |  |
|  | state | open, close |  | Yes |
| setRunStage |  | Mark the beginning of a new run stage; all statistics reported (e.g. temperature, current, etc) will have a timestamp relative to the time when the run stage began - setRunStage: { name: start the capture stage, runStage: all, frequency: 3 } |  |  |
|  | runStage | begin, capture, wash, end |  | Yes |
| setPosition |  | Sets the plunger of a cartridge syringe to either the fully retracted state (max) or fully compressed state (empty) - setPosition: { name: fully retract all cartridge syringe, socket: all, syringe: cartridge, position: max}} |  |  |
|  | socket | 0-7, all | all | No |
|  | pump | cartridge, manifold | cartridge | No |
|  | position | empty, max |  | Yes |
| setPump |  | Turn on or off a pump on one or more sockets at a particular speed. Pump will be turned off after duration elapses. - setPump: { name: start pulling cartridge pump at 800usec for all sockets, socket: all, pump: cartridge, speed: -800} |  |  |
|  | socket | 0-7, all | all | No |
|  | pump | manifold, cartridge |  | Yes |
|  | speed | rate to run motor at (uL/min); negative rates= pulling syringe; positive rates = pushing syringe |  | Yes |
| setTemperature |  | Tells an instrument socket to maintain the specified temperature (mode=auto) or PWM percent (mode=manual). You can optionally specify how quickly changes in temperature should be made by specifying the loopPeriod. Auto mode will attempt to run a PID controller. - setTemperature: { name: set socket 0 to maintain 5 degrees Celsius, socket: 0, mode: auto, setPoint: 5 } - setTemperature: { name: set all sockets to maintain 80% PWM, socket:all, mode: manual, setPoint: 80 } - setTemperature: { name: wait for up to 3 minutes for all sockets to stabilize 5C +/- 1C for 60 secs, duration: 180, socket:all, setPoint: 5, minStabilityTime: 60 } |  |  |
|  | socket | 0-7, all | all | No |
|  | mode | auto, manual | auto | No |
|  | setPoint | temperature in Celsius or PWM% |  | Yes |
|  | loopPeriod | response time in milliseconds between adjustments in temperature on the TEC module; shorter durations mean quicker responses to temperature variation from the set point. Minimum value = 200. | 4000 | No |
|  | minStabilityTime | minimum number of seconds required for temperature to stabilize at setPoint +/- maxTempDelta; if this parameter along with duration is specified, setTemperature will block until either the temperature stabilizes or duration elapses before a stable temperature can be achieved |  | No |
|  | maxTempDelta | the maximum variation in degrees C from setPoint which would be considered stable. Minimum value = | 1 | No |
| setValve |  | Turn a valve to a new position - setValve: { name: set manifold valve socket 0 to closed position, socket: 0, valve: manifold, position: closed } |  |  |
|  | socket | 0-7, all |  |  |
|  | valve | manifold, cartridge | cartridge | Yes |
|  | position | (For manifold valve): closed, primer, rinse, flush, elution, optional (For cartridge valve): manifold, sample, elution, closed |  |  |
| showDialog |  | Display a modal dialog on the instrument GUI and wait for a user to click a button - showDialog: { name: Are all 4 cartridges are inserted, labels: [{caption: Yes}, {caption: No, abort: true}] } |  |  |
|  | message | the text to display as the body of the dialog; if parameter not supplied, the step name will be used | -step name- | No |
|  | labels | List of button labels to use where each label is an object having the format: { caption: <caption-text>, abort: <true¦false> } | [ { caption: OK }, { caption: Cancel, abort: true} ] | No |
| sleep |  | Paused the method for a certain number of seconds - sleep: { name: wait for 5 seconds, duration: 5 } |  |  |
|  | duration | number of seconds to sleep |  | Yes |
| waitForBackgroundAction |  | Wait for last background action to complete. Always specify duration to set the maximum number of seconds to wait for the last executed background action to complete. - waitForBackgroundAction: { name: wait up to 120 seconds for pull fluid to complete, duration: 120 } |  |  |
| writeMemory |  | Write text to a particular memory address on the instrument - writeMemory: { name: write "test" to 100:100, address: "100:100", data: test } |  |  |
|  | address | address typically in hex to write to |  | Yes |
|  | value | text to write to the address above |  | Yes |

## Complete Example

method:

author: Kevin Haynie

version: 1.26

description: Test all actions

instruments: [exoprep,exodemo]

filmRequired: massSpec

reagentRequired: massSpec

minSampleVolume: 100

minEluateVolume: 50

minSyringeVolume: 300

manualMode: true

constants:

- DEPVoltage: 11

- DEPFreq: 2703

steps:

# showDialog

- showDialog: { name: this is a simple dialog with only an OK button }

- showDialog: { name: did the previous dialog show only an OK button?, labels: [{caption: Yes}, {caption: No, abort: true}] }

- showDialog: { name: this is a dialog with 3 custom buttons, labels: [{caption: button-1}, {caption: button-2 }, {caption: button-3 }] }

- showDialog: { name: test dialog, message: Did the previous dialog show 3 buttons?, labels: [{caption: Yes}, {caption: No, abort: true}] }

# playSound

- playSound: { name: play tada.mp3, filename: tada.mp3 }

- playSound: { name: play done.mp3 2 times, filename: done.mp3, numTimes: 2 }

- showDialog: { name: did you hear 3 audio files?, labels: [{caption: Yes}, {caption: No, abort: true}] }

# sendEmail

- sendEmail: { name: send test email to logged in user, subject: test email, body: this is a test - only a test }

- showDialog: { name: did you get the email?, labels: [{caption: Yes}, {caption: No, abort: true}] }

# reset

- reset: { name: reset noHoming, resetType: noHoming }

- sleep: {name: 'wait for reset to finish', duration: 3}

- showDialog: { name: did the instrument reset with no homing?, labels: [{caption: Yes}, {caption: No, abort: true}] }

- reset: { name: reset reHome, resetType: reHome }

- sleep: {name: 'wait for reset to finish', duration: 9}

- showDialog: { name: did the instrument reset with re homing?, labels: [{caption: Yes}, {caption: No, abort: true}] }

- reset: { name: reset atEase, resetType: atEase }

- sleep: {name: 'wait for reset to finish', duration: 6}

- showDialog: { name: did the instrument reset at ease?, labels: [{caption: Yes}, {caption: No, abort: true}] }

# prepare

# - prepare: { name: prepare start, prepareType: start }

# - prepare: { name: prepare priming, prepareType: priming }

# - showDialog: { name: did the preparation step for start and priming happen on the instrument?, labels: [{caption: Yes}, {caption: No, abort: true}] }

# setValve on manifold

- setValve: { name: set manifold valve socket 0 position closed, socket: 0, valve: manifold, position: closed }

- showDialog: { name: did the manifold valve for socket 0 close?, labels: [{caption: Yes}, {caption: No, abort: true}] }

- setValve: { name: set manifold valve socket 1 position primer, socket: 1, valve: manifold, position: primer }

- showDialog: { name: was the manifold valve for socket 1 set to primer?, labels: [{caption: Yes}, {caption: No, abort: true}] }

- setValve: { name: set manifold valve socket 2 position rinse, socket: 2, valve: manifold, position: rinse }

- showDialog: { name: did the manifold valve for socket 2 set to rinse?, labels: [{caption: Yes}, {caption: No, abort: true}] }

- setValve: { name: set manifold valve socket 3 position flush, socket: 3, valve: manifold, position: flush }

- showDialog: { name: did the manifold valve for socket 3 set to flush?, labels: [{caption: Yes}, {caption: No, abort: true}] }

- setValve: { name: set manifold valve socket 2 position elution, socket: 0, valve: manifold, position: elution }

- showDialog: { name: did the manifold valve for socket 0 set to elution?, labels: [{caption: Yes}, {caption: No, abort: true}] }

- setValve: { name: set manifold valve all sockets position closed, socket: all, valve: manifold, position: elution }

- showDialog: { name: did the manifold valve for all sockets close?, labels: [{caption: Yes}, {caption: No, abort: true}] }

# test when socket param not included - should default to all

- setValve: { name: set manifold valve all sockets position elution with no socket parm, valve: manifold, position: elution }

- showDialog: { name: did the manifold valve for all sockets set to elution?, labels: [{caption: Yes}, {caption: No, abort: true}] }

# setValve on cartridge

- setValve: { name: set cartridge valve socket 1 position manifold, socket: 1, valve: cartridge, position: manifold }

- showDialog: { name: did the cartridge valve for socket 1 set to manifold?, labels: [{caption: Yes}, {caption: No, abort: true}] }

- setValve: { name: set cartridge valve socket 2 position sample, socket: 2, valve: cartridge, position: sample }

- showDialog: { name: did the cartridge valve for socket 2 set to sample?, labels: [{caption: Yes}, {caption: No, abort: true}] }

- setValve: { name: set cartridge valve socket 0 position elution, socket: 0, valve: cartridge, position: elution }

- showDialog: { name: did the cartridge valve for socket 0 set to elution?, labels: [{caption: Yes}, {caption: No, abort: true}] }

- setValve: { name: set cartridge valve socket 3 position closed, socket: 3, valve: cartridge, position: closed }

- showDialog: { name: did the cartridge valve for socket close?, labels: [{caption: Yes}, {caption: No, abort: true}] }

# test when socket param not included - should default to all

- setValve: { name: set cartridge valve all sockets position closed with no socket parm, valve: cartridge, position: closed }

- showDialog: { name: did the cartridge valve for all sockets close?, labels: [{caption: Yes}, {caption: No, abort: true}] }

# setDEP

- setDEP: {name: turn DEP ON for all stages, stageId: all, voltage: $DEPVoltage, frequency: $DEPFreq}

- sleep: {name: 'sleep step', duration: 3}

- showDialog: { name: are all DEPs ON?, labels: [{caption: Yes}, {caption: No, abort: true}] }

- setDEP: {name: turn DEP OFF for all stages, stageId: all, voltage: 0, frequency: 0 }

- showDialog: { name: did all DEPs turn OFF?, labels: [{caption: Yes}, {caption: No, abort: true}] }

- setDEP: {name: turn DEP ON for stage 0, stageId: 0, voltage: $DEPVoltage, frequency: $DEPFreq }

- sleep: {name: 'sleep step', duration: 3}

- showDialog: { name: did DEPs for stage 0 turn ON?, labels: [{caption: Yes}, {caption: No, abort: true}] }

- setDEP: {name: turn DEP OFF for stage 0, stageId: 0, voltage: 0, frequency: 0 }

- showDialog: { name: did stage 0 DEPs turn OFF?, labels: [{caption: Yes}, {caption: No, abort: true}] }

- setDEP: {name: turn DEP ON for stage 1, stageId: 1, voltage: $DEPVoltage, frequency: $DEPFreq }

- sleep: {name: 'sleep step', duration: 3}

- showDialog: { name: did DEPs for stage 1 turn ON?, labels: [{caption: Yes}, {caption: No, abort: true}] }

- setDEP: {name: turn DEP OFF for stage 1, stageId: 1, voltage: 0, frequency: 0 }

- showDialog: { name: did DEPs for stage 1 turn OFF?, labels: [{caption: Yes}, {caption: No, abort: true}] }

# setPump

- setPump: {name: start pulling cartridge pump at 800usec for all sockets, socket: all, pump: cartridge, speed: -800}

- showDialog: { name: are all cartridge pumps pulling at 800usec/min right now?, labels: [{caption: Yes}, {caption: No, abort: true}] }

- setPump: {name: stop pulling cartridge pump at 800usec for all sockets, socket: all, pump: cartridge, speed: 0}

- showDialog: { name: did all cartridge pumps stop moving?, labels: [{caption: Yes}, {caption: No, abort: true}] }

- setPump: {name: start pushing cartridge pump at 800usec for all sockets, socket: all, pump: cartridge, speed: 800}

- showDialog: { name: are all cartridge pumps pushing at 800usec/min right now?, labels: [{caption: Yes}, {caption: No, abort: true}] }

- setPump: {name: stop pushing cartridge pump at 800usec for all sockets, socket: all, pump: cartridge, speed: 0}

- showDialog: { name: did all cartridge pumps stop moving?, labels: [{caption: Yes}, {caption: No, abort: true}] }

- setPump: {name: start pushing cartridge pump at 800usec for socket 2 for 5 seconds, socket: 2, pump: cartridge, speed: 800, duration: 5}

- showDialog: { name: did socket 2 cartridge pump push at 800 usec/min for 5 seconds?, labels: [{caption: Yes}, {caption: No, abort: true}] }

- setPump: {name: stop pushing cartridge pump at 800usec for socket 2, socket: 2, pump: cartridge, speed: 0}

- showDialog: { name: did socket 2 cartridge pump stop moving?, labels: [{caption: Yes}, {caption: No, abort: true}] }

# setEventTimer

- setEventTimer: { name: start event timer for all events every 3 seconds, stats: all, frequency: 3}

- sleep: {name: 'sleep for events', duration: 5}

- showDialog: { name: are both temp and current events arriving every 3 seconds?, labels: [{caption: Yes}, {caption: No, abort: true}] }

- setEventTimer: { name: stop event timer for all events every 3 seconds, stats: all, frequency: 0}

- showDialog: { name: did both temp and current events stop?, labels: [{caption: Yes}, {caption: No, abort: true}] }

# temp only reporting

- setEventTimer: { name: start event timer for temp events every 5 seconds, stats: temperature, frequency: 5}

- sleep: {name: 'sleep for events', duration: 7}

- showDialog: { name: are only temp events arriving every 5 seconds?, labels: [{caption: Yes}, {caption: No, abort: true}] }

- setEventTimer: { name: stop event timer for temp events every 5 seconds, stats: temperature, frequency: 0}

- showDialog: { name: did all events stop?, labels: [{caption: Yes}, {caption: No, abort: true}] }

# current only reporting

- setEventTimer: { name: start event timer for current events every 4 seconds, stats: current, frequency: 4}

- sleep: {name: 'sleep for events', duration: 6}

- showDialog: { name: are only current events arriving every 4 seconds?, labels: [{caption: Yes}, {caption: No, abort: true}] }

- setEventTimer: { name: stop event timer for current events every 4 seconds, stats: current, frequency: 0}

- showDialog: { name: did all events stop?, labels: [{caption: Yes}, {caption: No, abort: true}] }

# write memory

- writeMemory: { name: write "test" to 100:100, address: "100:100", value: test }

# setRunStage

- setRunStage: { name: set stage to begin, runStage: begin }

- sleep: {name: 'sleep in stage', duration: 2}

- showDialog: { name: is the current run stage set to begin?, labels: [{caption: Yes}, {caption: No, abort: true}] }

- setRunStage: { name: set stage to capture, runStage: capture }

- sleep: {name: 'sleep in stage', duration: 2}

- showDialog: { name: is the current run stage set to capture?, labels: [{caption: Yes}, {caption: No, abort: true}] }

- setRunStage: { name: set stage to wash, runStage: wash }

- sleep: {name: 'sleep in stage', duration: 2}

- showDialog: { name: is the current run stage set to wash?, labels: [{caption: Yes}, {caption: No, abort: true}] }

- setRunStage: { name: set stage to end, runStage: end }

- sleep: {name: 'sleep in stage', duration: 2}

- showDialog: { name: is the current run stage set to end?, labels: [{caption: Yes}, {caption: No, abort: true}] }

# setPosition

- setPosition: { name: fully retract all cartridge syringe, socket: all, syringe: cartridge, position: max}

- showDialog: { name: is the cartridge syringe fully retracted?, labels: [{caption: Yes}, {caption: No, abort: true}] }

- setPosition: { name: fully compress all cartridge syringe, socket: all, syringe: cartridge, position: empty}

- showDialog: { name: is the cartridge syringe fully compressed?, labels: [{caption: Yes}, {caption: No, abort: true}] }

- setPosition: { name: fully retract socket 0 of manifold syringe, socket: 0, syringe: manifold, position: max}

- showDialog: { name: is the manifold syringe on socket 0 fully retracted?, labels: [{caption: Yes}, {caption: No, abort: true}] }

- setPosition: { name: fully compress socket 0 of manifold syringe, socket: 0, syringe: manifold, position: empty}

- showDialog: { name: is the manifold syringe on socket 0 fully compressed?, labels: [{caption: Yes}, {caption: No, abort: true}] }

# setTemperature

- setTemperature: { name: set socket 0 to maintain 5 degrees Celsius, socket: 0, mode: auto, setPoint: 5 }

- showDialog: { name: is socket 0 temperature approaching and holding 5C?, labels: [{caption: Yes}, {caption: No, abort: true}] }

- setTemperature: { name: set all sockets to maintain 80% PWM, socket:all, mode: manual, setPoint: 80 }

- showDialog: { name: are all sockets approaching and holding 80% PWM?, labels: [{caption: Yes}, {caption: No, abort: true}] }

- setTemperature: { name: set all sockets to 10C, mode: auto, setPoint: 10 }

- setTemperature: { name: wait for up to 3 minutes for all sockets to stabilize 5C +/- 1C for 60 secs, duration: 180, socket:all, setPoint: 5, minStabilityTime: 60 }

- showDialog: { name: have all sockets stabilized at 5C +/- 1?, labels: [{caption: Yes}, {caption: No, abort: true}] }

showDialog: { name: are all sockets approaching and holding 10C?, labels: [{caption: Yes}, {caption: No, abort: true}] }

# getTECReport

- getTECReport: { name: refresh TEC information for all sockets }

- showDialog: { name: did the TEC socket information refresh?, labels: [{caption: Yes}, {caption: No, abort: true}] }